Standard Operating Procedure

Validation of Software and Computer Systems for ISO 17025

This is an example of a Standard Operating Procedure. It is a proposal and starting point only. The type and extent of documentation depends on the process environment. The proposed documentation should be adapted accordingly and should be based on individual risk assessments. There is no guarantee that this document will pass a regulatory inspection.

Publication from 
www.labcompliance.com

Global on-line resource for validation and compliance

Copyright by Labcompliance. This document may only be saved and viewed or printed for personal use. Users may not transmit or duplicate this document in whole or in part, in any medium. Additional copies and licenses for department, site or corporate use can be ordered from www.labcompliance.com/solutions.

While every effort has been made to ensure the accuracy of information contained in this document, Labcompliance accepts no responsibility for errors or omissions. No liability can be accepted in any way.

Labcompliance offers books, master plans, complete Quality Packages with validation procedures, scripts and examples, SOPs, publications, training and presentation material, user club membership with more than 500 downloads and audio/web seminars. For more information and ordering, visit www.labcompliance.com/solutions

| Company Name: |  |
| --- | --- |

| Controls: |  |
| --- | --- |
| Superseded Document | N/A, new |
| Reason for Revision | N/A |
| Effective Date | April 1, 2012 |

| Signatures: |  |
| --- | --- |
| Author | I indicate that I have authored or updated this SOP according to applicable business requirements and our company procedure: Preparing and Updating Standard Operating Procedures. Name: ________________________________ Signature: ________________________________ Date: ________________________________ |
| Approver | I indicate that I have reviewed this SOP, and find it meets all applicable business requirements and that it reflects the procedure described. I approve it for use. Name: ________________________________ Signature: ________________________________ Date: ________________________________ |
| Reviewer | I indicate that I have reviewed this SOP and find that it meets all applicable quality requirements and company standards. I approve it for use. Name: ________________________________ Signature: ________________________________ Date: ________________________________ |

- PURPOSE

Software and computer systems should be validated for compliance and business reasons. This SOP gives guidelines on how to validate commercial software and computer systems for IS0 17025.

- SCOPE

Validation of computer systems that have an impact on calibration and test results. Validation includes all life cycle phases from system planning to retirement. Exceptions to this procedure are possible but should be based on risk assessment and justified, documented and approved by laboratory management and QA. The SOP does not cover development activities and validation during development.

- GLOSSARY/DEFINITIONS

| Item | Explanation |
| --- | --- |
| Validation | Confirmation by examination and provision of objective evidence that particular requirements for a specific intended use are fulfilled. The degree of the validation needed depends on intended use. |
| Requirement specification: | The definition of what is required of a computing sys-tem in a specific intended use |
| Acceptance test: | Formal testing conducted to determine whether or not a computer system meets the requirement specification and to enable the laboratory to determine whether or not to accept the system. |
| GAMP® | Good Automated Manufacturing Practice (Forum). The GAMP Forum exists to promote the understanding of the regulation and use of computer and control systems within the pharmaceutical manufacturing industry. |
| GAMP® Category 3 | Standard software package. All applications problems are solved with standard functions. However, typically not all available functions are exercised by the user’s application. |
| GAMP® Category 4 | Configurable software package. Provides standard interfaces and functions that enable configuration of user specific applications. |
| GAMP Category 5 | Custom software package. Developed to meet specific needs of an application. Custom software may be a complete system or add on to a standard package. Custom software may be developed and supported in-house or by an external supplier. |
| Critical Requirement | Requirement that the user determines to be critical for the effective use of the system. |
| QA | Quality Assurance |

Note: For other definitions, see www.labcompliance.com/glossary.

- REFERENCE DOCUMENTS

  - GAMP, Good Automated Manufacturing Practice, A Risk-based Approach for Compliant GxP Computerized Systems , Version 5: 2008 (order from www.ispe.org).

  - SOP S-134: “Risk Assessment for Systems Used in GxP Environments”.
Available through www.labcompliance.com/solutions/sops.

  - SOP S-252: “Risk-Based Validation of Computer Systems”.
Available through www.labcompliance.com/solutions/sops.

  - SOP S-265: “Validation of Macro Programs and Other Application Software”.
Available through http://www.labcompliance.com/solutions/sops.

  - SOP S-274: “Quality Assessment of Software and Computer System Suppliers”.
Available through http://www.labcompliance.com/solutions/examples.

  - Template and Examples E-255: “Requirement Specifications for Chromatographic Data Systems”. 
Available through http://www.labcompliance.com/solutions/examples.

  - SOP S-262: “Change Control of Software and Computer Systems”. 
Available through http://www.labcompliance.com/solutions/sops.

  - SOP S-283: “Change Control for Networks and Systems - Planned Changes”.
Available through http://www.labcompliance.com/solutions/sops.

  - SOP S-284: “Change Control for Networks and Systems - Unplanned Changes”.
Available through http://www.labcompliance.com/solutions/sops.

  - Template and Examples E-362: “Test Case and Protocol – Authorized System Access”. 
Available through http://www.labcompliance.com/solutions/examples.

  - Template and Examples E-358: “Test Protocol For Excel Spreadsheet” (with traceability matrix, 29 pages).  
Available through http://www.labcompliance.com/solutions/examples.

  - Template and Examples E-326: “Network Infrastructure and System Identification”. 
Available through http://www.labcompliance.com/solutions/examples.

- RESPONSIBILITIES

  - Project owner

    - Owns the process to define, execute and document the validation activities and results. This requires the project owner to have experience in computer system validation.

    - Forms a validation project group.

    - Drafts validation documentation.

  - User Department

    - Provides inputs for requirement specifications.

    - Provides resources for testing.

    - Reviews and approves validation documents.

    - Advise on risk assessment

  - IT Department

    - Advises on risk assessment and the extent of testing related to network infrastructure.

    - Reviews and approves validation documentation related to network infrastructure.

    - Advises on Requirement Specifications related to ISO 17025 controls

  - Vendor

    - Provides functional specifications of the software and computer system.

    - Provides documented evidence that the software has been developed in a quality assurance environment and validated during development.

    - Accepts vendor audits, if necessary.

    - Provides information on how to prepare the site for installation of the computer system.

  - Plant Maintenance

    - Prepares site for installation of the computer system according to site preparation information provided by the supplier of the computer system.

  - Quality Assurance

    - Advises on requirements related to ISO 17025.

    - Reviews documentation for compliance with internal policies and ISO 17025.

    - Owns and conducts vendor assessments.

    - Reviews and approves validation documentation.

- FREQUENCY OF USE

  - Initially whenever computer systems are validated.

  - After system updates or any other changes to the system.

  - Whenever system reviews indicate that the system is in an out of validation state.

- VALIDATION PRINCIPLES

  - Overview

Validation of computer systems is not a once off event. For new systems it starts when a user department has a need for a new computer system and thinks about how the system can solve an existing problem. For an existing system it starts when the project owner gets the task of bringing the system into a validated state. Validation ends when the system is retired and all-important quality data is successfully migrated to the new system. Important steps in between are validation planning, defining user requirements, validation during development, vendor assessment for purchased systems, installation, initial and ongoing testing and change control. In other words, computer systems should be validated during the entire life of the system.

Because of the complexity and the long time span of computer validation the process is typically broken down into life cycle phases. An example is shown in the figure below.

User representatives define User or System Requirement Specifications (URS, SRS). The SRS or a special Request for Proposal (RFP) is sent to one or more vendors (see right side of the diagram). Vendors either respond to each requirement or with a set of functional specifications of a system that is most suitable for the user’s requirements. Users compare the vendor’s response with their own requirements. If none of the vendors meet all user requirements, the requirements may be adjusted to the best fit or additional software is written to fulfill the user requirements following the development cycle on the left side of the diagram. The vendor that best meets the user’s technical and business requirements is selected and qualified.

Next the system is installed, configured and well-documented. Before the system is used in a routine it should be tested in a suitable environment to verify functional specifications and in the final operating environment to meet user requirement specifications. Any change to the system should follow a documented change control procedure and before it is retired all quality and compliance relevant records generated on the system should be successfully migrated to the new system.

Activities for a specific validation project should follow a validation project plan. The plan outlines validation tasks, a time schedule, deliverables and owners for each deliverable. This validation project plan is derived from a company or a site validation master plan. Validation summary results are documented in a validation report.

  - Software Categories

The extent of validation depends on the type of software, the complexity of the computer system and on the risk or impact a computer system has on calibration or test results. The extent of validation at the user’s site also depends on the widespread use of the same software product and version. The more a specific software is used and the less customization made for a specific software the less testing is required by individual users. GAMP (Ref4.1 ) has developed software categories based on the level of customization. In total there are five categories. Category one are operating systems and category two is firmware that controls automated instruments. Both categories don’t require separate validation because they are either validated as part of the application software (Category 1) or as part of equipment qualification (Category 2). Therefore, in the context of this SOP only categories three to five are of interest. Definitions can be found under Glossary/Definitions in section 3. Each computer system should be associated to one of the three categories.

| Category | Description |
| --- | --- |
| GAMP 3 | Standard software package. No customization. Examples: MS Word (without VBA scripts). Computer controlled spectrophotometers. |
| GAMP 4 | Standard software package. Customization of configuration. Examples: LIMS, Excel spreadsheet application where formulae and/or input data are linked to specific cells. Networked data systems. |
| GAMP 5 | Custom software package. Either all software or a part or the complete package has been developed for a specific user and application. Examples: Add-ons to GAMP Categories 3 and 4, Excel with VBA scripts. Custom built software and systems. |

  - Risk Assessment

The extent of validation also depends on the risk that the records generated or processed by the system have on product quality. Therefore risk categories should be defined for each system. The risk category of the system depends on the risk levels and number of critical records processed by the system. Typically risk categories are defined as high, medium and low.

- PROCEDURE

  - Proposal and Planning

    - A user representative requests a new computer system using the form in Attachment 9.1. The request should include information on the intended use of the system, the intended location and environment, how the problem is currently solved and a short description of the suggested new system. The request should also include business benefits, cost estimates and a list of possible suppliers.

    - The request is sent to the laboratory manager and IT for review and approval

    - If the request is approved by the laboratory manager and IT, proceed to 8.1.4 otherwise stop here.

    - The laboratory manager designates a project owner.

    - The project owner forms a validation group consisting of representatives from:

- Anticipated users of the system.

- Quality Assurance (QA).

- IT department (if the computer system is planned to be networked).

  - The project owner prepares a first draft of the validation project plan and distributes the plan for review by the validation group.

The project plan should include an initial risk assessment of the system.

  - Setting Specifications

    - Project owner drafts a user requirement specifications document based on inputs from:

- Anticipated users of the system to address technical requirements.

- Laboratory manager to address business requirements.

- IT department

- QA department to address quality standard and internal policy requirements.

Special considerations should be given to:

- Description of the intended process and environment.

- Functions important for executing critical steps.

- Functions that are required by standards, e.g. ISO 17025 or by regulations, such as FDA’s GMP or 21 CFR Part 11.

- Security functions.

- Functions to ensure data integrity, e.g., electronic audit trail

- Compatibility with current and future network environments.

- Upgradeability for future applications.

- Documented evidence from the supplier for validation during development in a quality assurance environment and willingness to accept vendor audits.

- Services supplied by the supplier, e.g., familiarization, training, installation qualification, operational qualification and ongoing support (phone, on-site) with desired response time.

- Testability of functions.

- Unique identification of all functions, e.g., through numbers.

- Functions can have priorities, e.g. must, want or nice to have.

- Consecutive validation activities, e.g., operational qualification and performance qualification tests should be traceable back to users requirements.

    - The project owner distributes the draft requirement specifications document to the input team in 8.3.1 for review, collects inputs and updates the document, if necessary.

    - The project owner pre-selects suppliers and sends the requirement specifications document or a special Request for Proposal (RFP) to the pre-selected suppliers. The RFP is derived from the requirement specifications.

    - The project owner evaluates the responses of the vendors and makes a proposal to the team as identified in 8.3.1 on which vendor should be selected. The decision should be based on:

- Meeting requirement specifications as defined in 8.3.2.

- Business experience with the vendor.

- Costs for purchasing and implementation.

    - The project owner identifies gaps between the requirement specifications and the functionality as provided by the vendor’s system. If there are no gaps proceed to 8.4.

    - The project owner presents the results from 8.3.5 in a meeting with the input team as identified in 8.3.1.

    - The team decides if the gaps are acceptable or if missing functionality must be added through additional customized software.

    - If the gaps are acceptable, proceed to 8.3.10, if not proceed to 8.3.9.

    - New or additional software should be developed to close the gap between the user’s requirement specifications and the standard software. Such software belongs to GAMP Category 5. For validation during development follow the SOP in Reference 4.4.

    - The project owner updates the project validation plan with more details on which approach to take for extent of validation and testing based on risk assessment of the system, user requirements and GAMP Category.

  - Vendor Assessment

    - QA department performs a vendor assessment following the SOP in Reference 4.5. Procedures for assessment can be:

- Using and documenting external references.

- Using and documenting internal references and experience.

- Mail audit through checklists and follow-up by phone.

- 3rd party audit.

- Direct audit.

The final decision on which procedure to follow depends on:

- Complexity of the system.

- Business impact of the system.

- Maturity or widespread use of the system.

    - QA identifies a vendor as “acceptable” or “not acceptable” and documents details of the vendor assessment.

    - If the vendor is not acceptable as a supplier of the software, go back to 8.3.3.

  - Installation and Configuration

    - The project owner requests information from the vendor on environmental conditions and space requirements. Such information should include:

- Temperature and humidity.

- Electromagnetic interference.

- Power supply.

- Space requirements.

- For networked systems: network infrastructure.

    - Project owner makes arrangements with plant maintenance (and with IT if the computer system is networked) to prepare the site for installation of the computer system.

    - Upon arrival of the system the project owner verifies completeness according to the purchase order and checks the physical condition of the system.

    - The project owner arranges installation of the system with the vendor.

    - A vendor or user representative installs and configures the system according to the vendor’s recommendations.

    - The vendor or user representative verifies adequate installation of hardware and software. Verification activities can include:

- Inspection of hardware installations.

- Verification of accurate and complete software installation.

- Execution of tests as defined in the vendor’s installation and start up instructions.

    - The project owner documents the system and all configuration settings using the form in Reference 4.11 or equivalent or special software designed for configuration management.

    - The project owner checks if back-up copies of software are available or arranges for making back-up copies according to national and international copyright regulations. The project owner arranges for storage of back-up copies in a safe environment.

    - Installation protocols are signed off by the user firm’s QA department,  the project owner and vendor representatives if the installation is carried out by the vendor.

  - Functional Testing in a Suitable Environment

These tests should ensure that the system performs all functions as intended in a suitable environment.

    - A vendor or user representative verifies adequate functioning of the computer system. Special attention should be given to testing of:

- Business critical functions.

- All non-standard functions. This can mean testing all functions if the software has been developed specifically for an individual user firm.

- Physical and logical security.

- Back-up and retrieval of data files.

- Standard functions without evidence from the vendor that the correct functioning has been verified in a suitable environment.

- Functions that can be influenced by user specific configuration settings.

- Functions that can be influenced by the user’s environment, e.g., distance between networked components, data traffic and operator skills.

- Start-up and shutdown of applications.

- Error handling.

- Parallel applications.

- For networked systems: 
- Connectivity of network components.
- Accuracy of file transfer.

Considerations for testing:

- Tests should be linked to user requirements.

- Tests should be quantitative, which means that test protocols should include acceptance criteria and actual results.

- Test cases and data files should be defined and documented for reuse.

- Tests should include test data across the entire application range, boundaries and data outside allowed ranges.

An example for a simple test is shown in Reference 4.10. An example with detailed test scripts, boundary testing, out of range testing and a traceability matrix is shown in Reference 4.11.

The nature and extent of the testing depends on the criticality, maturity, complexity and level of customization of the system. For the nature and extent of testing and documentation based on the above criteria follow the SOP in Reference 4.3.

    - Test protocols are signed off by the user’s QA department, the project owner and a vendor representative if testing is performed by the vendor.

    - Tests as defined in 8.6.1 should be regularly repeated. Test frequency is once per year.

  - Initial and Ongoing System Performance Checks

This test should ensure that the user’s application performs as intended in the user’s environment.

- A user representative verifies adequate functioning of the user’s application in the user’s environment. For extent of testing and documentation follow the SOP in Reference 4.3.

    - Test protocols are signed off by the QA department and the project owner.

    - Tests in 8.7 should be repeated during routine use of the system. Type of tests and frequency of testing depends on the type and criticality of the system and should be defined by the project owner.

    - To ensure smooth ongoing performance users of the system should perform additional tasks:

- Regular virus checks.

- Regular removal of temporary files.

- Regular system access checks.

    - QA should perform a periodic review of the computer system. Frequency should be yearly.  For initial and ongoing reviews the checklist in Reference 4.11 can be used.

  - Change Control

    - Any change to the system should follow a change control procedure:

- Use SOP in Reference 4.7 for systems not networked.

- Use SOP in Reference 4.8 for networked systems if the change is planned.

- Use SOP in Reference 4.9 for networked systems if the change is not planned.

Special considerations should be given to:

- Risk assessment to estimate the business impact of the change.

- An assessment of the change impact on the validation state to estimate efforts for revalidation.

- Decision who can initiate, authorize and approve changes for different system categories.

- Retrospective approval of unplanned changes.

    - All changes should be recorded in change history logs using templates as provided in parts of the SOPs defined in section 8.8.1 or equivalent.

  - Retirement

    - The project owner initiates a retirement process using the form in 9.2.

Special consideration should be given to the question whether critical quality and/or compliance data have been processed on the system and if these data need to be reprocessed after retirement.

    - If data do not need to be reprocessed after retirement proceed to 8.9.4.

    - The project owner arranges for migration of the critical records to the new system and for verification of correct reprocessing of the data on the new system.

    - The project owner arranges for conversion of non-critical records to standard electronic formats.  Non-critical records can be converted to standard formats such as PDF or XML or can be printed. The content and meaning of the records should be preserved.

    - Project owner distributes the retirement form to QA and operation’s management (and for networked systems to the IT department) for review and approval.

    - Project owner updates the system inventory database.

  - Validation Report

    - The project owner drafts a validation report after the first PQ tests in Reference 8.7 are completed. The report should include:

- Brief description of the system.

- Validation approach.

- Summary of test results.

- Any deviations from the validation plan and corrective actions.

- Statement that the system is validated.

- Sections for approvals by QA and operation’s management and by the IT department if the system is networked.

    - The project owner distributes the initial validation report to QA and operation’s management, and to the IT department if the system is networked, for approval.

  - Validation of Existing Systems

There may exist a need to validate existing systems when:

  - Systems are used in ISO 17025 environments but have not been validated according to current standard requirements.

  - Systems currently not used for ISO 17025 or critical business environments will be used in the future in GxP environments or for critical business applications.

In general procedures to validate existing systems are not significantly different from validating new systems. Steps from 8.6 to 8.10 are the same.

Steps 8.1 to 8.5.9 are replaced by 8.11.1 to 8.11.4.

    - Project owner forms a validation group and drafts a validation project plan.

    - Project owner documents current use of the system and system functionality.

    - The project owner evaluates if current functionality meets all critical business and regulatory requirements. Examples are:

- Security functions.

- Functions as required by the ISO 17025.

    - The project owner presents the results to the validation group. If the group decides: 
a) The system can be used as is, proceed to 8.5.7.
b) The system can be used after upgrades, proceed to 8.3.1.
c) The system is not suitable even after an upgrade, stop the procedure.

- ATTACHMENTS

  - Attachment – Computer System Request Form

| Requester: |  |
| --- | --- |
| Intended Use of the System: |  |
| Intended Location and Environment: |  |
| How the Problem is Solved at Present: |  |
| Description of the New System and Business Benefits: |  |
| Possible Suppliers: |  |
| Estimated Costs: |  |

  - Attachment – System Retirement Request Form

| Project owner: |  |
| --- | --- |
| System ID: |  |
| Location of the System: |  |
| Reason for the Retirement: |  |
| Anticipated Time Frame for Retirement: |  |
| Have critical quality or compliance data been processed on the system? If yes, describe type of data: |  |
| Is there a business or compliance need to reprocess the data after system retirement? If yes, justify and document: |  |
| Proposal for Retirement Process: |  |
| Approvals and Signatures: |  |
| Quality Assurance | ______ ____________ ____________ Date Printed Name Signature |
| Operation’s Manager | ______ ____________ ____________ Date Printed Name Signature |
| Project owner | ______ ____________ ____________ Date Printed Name Signature |